## Slide 1: SEMI-AUTOMATIC QUALITY ASSURANCE (SAQA) MATLAB SCRIPT FOR ACR MRI PHANTOM
Jidi Sun1,2,3, M. Barnes2, P. Stanwell1, J. Dowling3, F. Menk1, P.B. Greer1,2
1 University of Newcastle, Newcastle, New South Wales, Australia.
2 Calvary Mater Newcastle, Newcastle, New South Wales, Australia.
3 CSIRO Australian E-Health Centre, Brisbane, Queensland, Australia.
EPSM2013, Perth

## Slide 2: Content
ACR MRI phantom background
SAQA Matlab Script
Manual QA vs. SAQA
How to make a better phantom

## Slide 3: What is the ACR MR Phantom?
Localiser
S 1
S 2
S 3
S 4
S 5
S 6
S 7
S 8
S 9
S 10
S 11
American College of Radiology (ACR)
  148 cm length, 190 cm diameter
  10 mmol NiCl2, 45 mmol NaCl
  T1 & T2 images

## Slide 4: MR QA with ACR
Localiser
S 1
S 2
S 3
S 4
S 5
S 6
S 7
S 8
S 9
S 10
S 11
Geometric distortion
High contrast spatial resolution
Slice thickness accuracy
Slice position accuracy
Image intensity uniformity
Ghosting
Low contrast object detectability
[B0 homogeneity, field drifting, RF coil check, inter-slice RF interference]

## Slide 5: Well Documented

## Slide 6: Available Software
http://www.jwpacific.com/radiadiag.html
http://www.onpointmd.com/index.php?option=com_content&view=article&id=47&Itemid=55

## Slide 7: What does SAQA aim to achieve?
Free open-source script
Replace all the manual image window/level adjustment, length measurements, ROI generation, ROI area calculation, mean intensity calculation
Reduce error that is caused by manual process
Reduce physicist QA time
Source code possibly be implemented into other QA phantoms

## Slide 8: What is SAQA?
SAQA: semi-automatic QA

## Slide 9: Image Thres. (Window/Level)
Thres. image
Orig. image
Slice position accuracy
Geometric distortion
Slice thickness accuracy

## Slide 10: ROI Creation
Areaextreme=1 cm2
Areawater=195-205 cm2
Areaair=10 cm2
ACR recommended area:
Uniformity Test
Ghosting Test

## Slide 11: Manual Process (Calvary Mater Hospital)
Manual QA initially made:
  ImageJ, NIH, USA
    Visualise DICOM image slices
    Manual thresholding image
    Manual measurements made using ruler and ROI tools
  Excel
    Documenting and report

## Slide 12: Method: Manual vs. SAQA
Time for each method estimated
ACR criteria agreement
Manual vs. SAQA difference

## Slide 13: Time Conservation
Time:
  Manual: 45 mins + writing Excel report
  SAQA: 5 mins, report generated automatically

## Slide 14: ACR Criteria Agreement
a
c
b

## Slide 15: What can we do to make a better phantom?
Make sure the uniformity test slice is away from any inserted structure, for example one slice separation (e.g. S7)
Make critical structure away from the phantom edge (reduce the Gibb’s artefact, e.g. low contrast disk radius 3)
Make use of the grid on S5

## Slide 16: What can we do to make a better phantom?
Make sure the uniformity test slice is away from any inserted structure, for example one slice separation (e.g. S7)
Make critical structure away from the phantom edge (reduce the Gibb’s artefact, e.g. low contrast disk radius 3)
Make use of the grid on S5

## Slide 17: Uniformity Slice
S7
S7
S6
S6
S5
S8

## Slide 18: What’s wrong?
Problem:
Altered signal at the water and structure interface  underestimate PIU
Possible causes:
Slice positioning error (most depended factor)
Magnetic susceptibility artefact
Solution:
Increase the phantom length, so that a slice exists between the UNIFORMITY test slice and the nearby slices with inserted structures  (Slice 6.5)
S5
S6
S6.5
S7
S8

## Slide 19: What can we do to make a better phantom?
Make sure the uniformity test slice is away from any inserted structure, for example one slice separation (e.g. S7)
Make critical structure away from the phantom edge (reduce the Gibb’s artefact, e.g. low contrast disk radius 3)
Make use of the grid on S5

## Slide 20: Edge Effect on Contrast Result

## Slide 21: What can we do to make a better phantom?
Make sure the uniformity test slice is away from any inserted structure, for example one slice separation (e.g. S7)
Make critical structure away from the phantom edge (reduce the Gibb’s artefact, e.g. low contrast disk radius 3)
Make use of the grid on S5

## Slide 22: Geometric Distortion on S5
Implement grid cross-sections identification script from our pelvic phantom work last year into ACR MRI phantom to get the 2D distortion information

## Slide 23: Conclusion
Most manual QA processes have been replaced by SAQA script
SAQA reduces QA time (45 mins 5 mins)
SAQA results comparable to manual results
Suggestions on making better QA phantom learned from ACR MRI phantom
Limitation of SAQA: high intensity external marker
Possibility to implement SAQA for other QA phantoms

## Slide 24: Acknowledgement
Michael Barnes [MP], closely working together
Peter Stanwell [MR Physicist]
Jameen Arm, Leah Best, Melissa Murphy and Nick Marks [Radiographer]
Seeded region growth script written by Stéphane (http://www.mathworks.com.au/matlabcentral/fileexchange/35269-simple-single-seeded-region-growing) – used for low contrast disk masking

## Slide 25: Thank you! 